## Itron
Bit | Display | Description | 403 | 603 | 803
0 | 10000000 | Power off | x | x | x
1 | 20000000 | Low battery | x | x | x
2 | 90000000 | External alarm (i.e. via KMP) | x | x | x
3 | 1000000 | t1 above range or disconnected | x | x | x
4 | 100000 | t2 above range or disconnected | x | x | x
5 | 2000000 | t1 below range or short circuited | x | x | x
6 | 200000 | t2 below range or short circuited | x | x | x
7 | 9900000 | Wrong dt (t1-t2) | x | x | x
8 | 3000 | V1 air | x | x | x
9 | 4000 | V1 reverse | x | x | x
10 | 5000 | V1 low signal | x | x | x
11 | 6000 | V1 > qs for more than 1 hour | x | x | x
12 | 80 | InA water system leak | x | x | x
13 | 8 | InB water system leak | x | x | x
14 | 90 | InA external alarm | x | x | x
15 | 9 | InB external alarm | x | x | x
16 | 1000 | V1 com. error |  | x | x
17 | 2000 | V1 pulse error |  | x | x
18 | 70 | InA2 water system leak |  | x | x
19 | 7 | InB2 water system leak |  | x | x
20 | 10000 | t3 or t4 above range or disconnected |  | x | x
21 | 20000 | t3 or t4 below range or short circuit |  | x | x
22 | 100 | V2 com. error |  | x | x
23 | 200 | V2 pulse error |  | x | x
24 | 300 | V2 air |  | x | x
25 | 400 | V2 reverse |  | x | x
26 | 500 | V2 low signal |  | x | x
27 | 600 | V2 > qs for more than 1 hour |  | x | x
28 | 7000 | V1V2 Burst Out |  | x | x
29 | 700 | V1V2 Burst In |  | x | x
30 | 8000 | V1V2 Leak Out |  | x | x
31 | 800 | V1V2 Leak In |  | x | x

## Diehl Sharky
Bit | Description | usage
0 | reserved | -
1 | any application error | -
2 | power low | E-8, E-9
3 | permanent error | C-1, E-4
4 | temporary error | E-1, E-3, E-6, E-7, leak error
5 | manufacturer specific | see table below
6 | manufacturer specific | see table below
7 | manufacturer specific | see table below
Error | C-1 | E-8 | E-4 | E-1 | E-7 | E-9 | E-3 | E-6 | Leak error | E-5
Mbus status Byte | 0x08 | 0x04 | 0x28 | 0x50 | 0x70 | 0x84 | 0xB0 | 0xD0 | 0xF0 | 0x10
Prio | Highest | [8 empty] | Lowest
Error display | Meaning
C-1 | Basic parameter error in flash or RAM
E-1 | Temperature measurement error
Temperature range exceeded [-9.9 °C … 190 °C]
Sensor short-circuit
Sensor break
E-3** | Temperature sensors reversed in hot and cold lines
E-4 | Hardware error in ultrasonic measurement
Ultrasonic transducer defective
Short-circuit in ultrasonic transducer
E-5 | Reading too frequently
M-Bus communication not possible for short time
E-6** | Wrong direction of flow
Flow sensor incorrectly installed
E-7 | No meaningful ultrasonic receive signal
Air in the measuring path
E-8 | No primary voltage (only if mains unit used)
Powered by back-up battery
E-9 | Warning: battery nearly exhausted
E-A* | Leakage: pipe break detected
E-b* | Leakage: leakage detected in energy meter
E-C* | Leakage: leakage pulse input 1
E-d* | Leakage: leakage pulse input 2

## UH50
Bit-No | Decimal value | Identifier on meter | Explanation
0 | 1 | F0 | Error during flow metering (e.g. Air in measuring pipe)
1 | 2 | F1 | Interruption of flow temperature sensor
2 | 4 | F2 | Interruption of return temperature sensor
3 | 8 | F3 | Electronic for temperature evaluation defective
4 | 16 | F4 | Battery empty 1
5 | 32 | F5 | Short-circuit flow temperature sensor
6 | 64 | F6 | Short-circuit return temperature sensor
7 | 128 | F7 | Fault in the internal memory (CRC)
8 | 256 | F8 | Error F1, F2, F5 or F6 pending for longer than 8h.
9 | 512 | F9 | Error in the electronics
10 | 1024 | F0V | Prewarning for soiling of the measurement tube
11 | 2048 | F7V | Correctable error in the internal memory EEPROM 2
12 | - | - | Always 0
13 | - | - | Always 0
14 | - | - | Always 0
15 | - | - | Always 0

## UH30
Bits | Meaning
ZZ.7 | Time error DSMR
ZZ.6 | Leakage warning (water meters)
A.3 | 0 = Installation in return flow position
1 = Installation in forward flow position
A.1 | Incorrect flow direction (FIow-Negativ)
A.0 | Negative temperature difference (Difference-Negative)
B.3 | Installation error volume measurement part
B.2 | Installation error sensor
B.1 | F4 pre-warning
B.0 | F0 pre-warning
C.0 | 0 = Installation location cannot be changed when calibration seal is set
1 = Installation location can be changed when calibration seal is set
D.3 | 0 = Amount of energy in case of incorrect installation
1 = Amount of cooling energy
D.2 | Reserved
D.1 | Fault in the electronics (F9)
D.0 | Errors F1, F2, F3, F5 or F6 for longer than 8 hours, recognition of attempts to manipulate. No further measurements are carried out. (F8)
E.3 | Fault in internal memory holding (EEPROM) (F7)
E.2 | Short-circuit return flow temperature sensor (F6)
E.1 | Short-circuit forward flow temperature sensor (F5)
E.0 | Problem with the power supply; Battery empty (F4)
F.3 | Electronics for temperature evaluation defective (F3)
F.2 | Interruption in the cold side temperature sensor (F2)
F.1 | Interruption in the hot side temperature sensor (F1)
F.0 | No flow can be measured (F0)